## Risk Assessment - Master Sheet
Activity-Based Risk Assessment Form
| Name of Department | Singapore Synchrotron Light Source |  | Location of Lab |  | 5, Research Link, Singapore 117603
| Name of Laboratory | SUV Beamline |  | Name of PI |  | Prof. Mark B H Breese
| Name of Researcher/LO | Luis Rodriguez/Edwin Ong |  | Name of Activity/Experiment |  | Thin film growth on PLD system
No | Description/Details of Steps in Activity | Hazards | Possible Accident / Ill Health & Persons-at-Risk | Existing Risk Control (Mitigation) | Severity | Likelihood (Probability) | Risk Level | Additional Risk Control | Person Responsible | By (Date)
Important! Users must wear Laser safety goggles with the following minimum optical densities (OD): 1064 nm --> OD=6.1, 532 nm --> OD=6.8, 355 nm --> OD=2.2, and 266 nm --> OD=3.7. There is no interlock for the black casing enclosing the laser path within the system, Class 4 laser can still operate without the black casing. Therefore, do not power up Class 4 laser without the black casing at all time.
1 | Sample preparation: Mount sample onto Neocera sample holder and load sample into Load Lock Chamber of PLD system. | Nil | Nil | Nil
2 | Pump down Load Lock Chamber via computer control | Nil | Nil | Nil
3 | Open actuator valve between Load Lock Chamber and main chamber of PLD system. Lift transfer arm up to sample position height for transfer | Physical hazard - Dropping of transfer arm probe during transfer | Physical injuries (bruises) if transfer arm probe is dropped onto hand. | 1) Lift transfer arm with care (i.e. both hands if necessary) 2) Secure transfer arm properly with secure-hooks (supplied by PLD system) or with the screw of the transfer arm probe to prevent arm probe from moving | 1 | 1 | 1
4 | a. Transfer sample from Load Lock Chamber transfer arm to main chamber transfer arm. b. Initiate manipulator position for sample transfer (Z » 40 mm) using the PLD software. Note that PLD software should already be running on the PC, c. Transfer sample from main chamber transfer arm to main chamber manipulator. d. Initiate manipulator position to growth position (Z » 86mm). | Nil | Nil | Nil
5 | Open valve between RHEED gun and main chamber. Open RHEED screen shutter. Turn on RHEED gun power supply. Adjust RHEED settings to acquire RHEED pattern. | Electrical hazard from operating high voltage RHEED gun power supply (>20kV) | Electric shock or burns | 1) Inspect RHEED gun power supply before operation 2) Follow instruction manual on the steps to operate the RHEED gun properly 3) Do no operate power supply with wet-hands | 2 | 1 | 2
|  | Radiation hazard from stray x-ray generated from electron beam of RHEED gun inside the chamber. | X-ray irradiation damage to eye and body. | Ensure all viewports are covered with lead-glass supplied by Neocera. Annual inspection for radiation using radiation survey meter to monitor for radiation leak. All personnel must have radiation worker license and wear radiation dosage badge when working. Personnel must complete relevant NUS radiation safety training and equipment training | 3 | 1 | 3
6 | Turn on valves of gas tanks for desired gas (O2, N2 and/or Ar). Initiate Mass Flow Control (MFC) for desired gases required for growth - O2, N2 or Ar. Flow is controlled using the PLD software. | Physical hazard - Gas cylinder toppling onto personnel | May cause bruises or factures | Ensure all gas tanks secured in gas cylinder stand. | 3 | 1 | 3
|  | Fire hazard from excessive O2 gas inside chamber. Fire resulting from ignition of O2 gas | Ignition of O2 gas inside chamber may cause viewports that are covered with lead-glass to crack and explode. Lead glass fragments fly out and cause lacerations to nearby personnel | Manufacturer ensure that system is safe for use at O2 gas pressure = 500 Torr and Temperature = 1000 degrees. Ensure that users are properly trained and follow Safe Work Procedure before using system | 2 | 1 | 2
|  | Fire hazard from O2 leakage due to poor installation of gas regulator. Fire resulting from ignition of O2 gas | May cause burns | Ensure that O2 tank has no leak by listening for hissing sound at the gas regulator. If there is a hissing sound, it means that O2 gas is leaking. Bubble or soap test to ensure no leak at the connections. | 2 | 1 | 2
7 | Initiate sample heating using sample heating profile from software | Laser hazard from laser-heater (class 4 laser) shine into eyes | Eye injury from laser | 1) PLD system laser heater is fully enclosed to prevent stray radiation from laser heater. 2) Only trained personnel with appropriate laser licence (N3) can operate the laser of the system. 3) Ensure sample is seated in manipulator before heating. 4) Do not start laser heater if there is no sample in the manipulator. 5) Personal Protective Equipment must be worn: laser safety goggles | 3 | 1 | 3
8 | Turn on PLD pulsed laser power supply and cooling system | Electrical hazard from operation of electrical equipment | Electric shock or burns | 1) Inspect power supply and cooling system before operation 2) Do not operate supply and cooling system with wet-hands | 2 | 1 | 2
9 | Activate pulsed laser for pre-ablation of target. Activate pulsed laser for desired number of pulses for thin film growth. | Laser hazard from laser-heater (class 4 laser) shine into eyes | Eye injury from laser | 1) PLD system pulsed laser system is fully enclosed - no stray laser from laser heater will be observed from outside the system. 2) Only trained personnel with appropriate laser licence (N3) can operate the laser of the system. 3) Ensure all viewports are properly covered with laser optical filters before starting the pulsed laser. 4) Do not start pulsed laser if there is no target in the chamber. 5) Personal protective Equipment must be worn: laser safety goggles | 2 | 1 | 2
10 | Initiate gas backfill process using the software. | Fire hazard from excessive O2 gas inside chamber. Fire resulting from ignition of O2 gas | Ignition of O2 gas inside chamber may cause viewports that are covered with lead-glass to crack and explode. Lead glass fragments fly out and cause lacerations to nearby personnel | Manufacturer ensure that system is safe for use at O2 gas pressure = 500 Torr and Temperature = 1000 degrees. Ensure that users are properly trained and follow Safe Work Procedure before using system | 3 | 1 | 3
11 | After the heating process has been completed, wait for sample is sufficiently cooled before transferring to Load Lock chamber. | Direct skin contact with sample while it is still hot | May cause minor burns | Ensure that users are properly trained and follow Safe Work Procedure (SWP-SSLS-24-RSXS PLD System) before using system. | 1 | 1 | 1
12 | Pump down chamber to remove the O2 gas from the chamber. Transfer sample from process chamber to Load Lock chamber. Retrieve sample from Load Lock chamber. | Nil | Nil | Nil
13 | Turn off MFC for gas-flow into chamber. Turn off gas valves from gas tanks. | Physical hazard - Gas cylinder toppling onto personnel | May cause bruises or factures | Ensure all gas tanks secured in gas cylinder stand. | 3 | 1 | 3
|  | Fire hazard from O2 leakage due to poor installation of gas regulator. Fire resulting from ignition of O2 gas | May cause burns | Ensure that O2 tank has no leak by listening for hissing sound at the gas regulator. If there is a hissing sound, it means that O2 gas is leaking. Bubble or soap test to ensure no leak at the connections. | 2 | 1 | 2
14 | Shut-down RHEED power supply, pulsed laser power supply and cooling system. | Electrical hazard from operation of electrical equipment. | Electric shock or burns | 1) Inspect power supply and cooling system before operation. 2) Do not operate supply and cooling system with wet-hands. | 2 | 1 | 2
| Conducted By | Luis Rodriguez |  | Approved By
|  | Edwin Ong |  | Name | Prof. Mark B H Breese
Signature
|  |  |  | Approval date |  | 2018-07-31 |  | Next Revision date | 2021-07-31
(Maximum 3 years)
Note on C3: Write Name of Your department
Note on G3: Location of Lab e.g S9 # 05 - 04
Note on C5: Write Name of Your department
Note on G5: Name of your professor/ guide
Note on G7: Name of the experiment e.g.. Distillation of THF
Note on G9: NUS: : (after consideration of mitigation)
Note on I36: NUS: At least once every three (3) years or; When there is changes to work process or; Incident / Accident

## Risk Ranking - Guide
|  | Likelihood |  |  | Risk = Likelihood x Severity
|  | Likely | Possibly | Unlikely | RISK | DECISION PROCESS
Severity | Low | 3 | 2 | 1 | < 3 | RISK ACCEPTABLE
| Med | 6 | 4 | 2 | 3, 4 | CONSIDER ADDITIONAL RISK CONTROL
| High | 9 | 6 | 3 | > 4 | ADDITIONAL RISK CONTROL REQUIRED
Likelihood
| 1 | Unlikely |  | Not likely to occur (has not occurred in the PI's Lab or similar Lab setup.)
| 2 | Possible |  | Possible or known to occur (has occurred in the PI's Lab or Similar Lab setup.)
| 3 | Very Likely |  | Common or repeating occurrence (has occurred repetitively in the PI's Lab or similar Lab setup.)
Severity
| 1 | Low |  | (e.g. No injury, injury or ill-health requiring first aid treatment only - includes minor cuts and bruises, irritation, ill-health with temporary discomfort)
| 2 | Medium |  | (e.g. Injury requiring medical treatment or ill-health leading to disability – includes lacerations, burns, sprains, minor fractures, dermatitis, deafness, work-related upper limb disorders)
| 3 | High |  | (e.g. Fatal, serious injury or life-treatening occupational disease – includes amputations, major fractures, multiple injuries, occupational cancer, acute poisoning and fatal diseases)
Severity - Consider the magnitude/severity of the consequences of the Risk Factor occurring and then list this as 3 (High), 2 (Moderate) or 1 (Low). Severity normally will not change unless there is a physical change to the equipment or process.
Likelihood - Team should rely upon their experience and consider realistic scenarios. Listed below are examples of factors that may be considered in determining the likelihood. - Past experience / incidents - Complexity of the activity - Number of personnel involved in the activity (e.g. all personnel, a limited number of trained personnel, etc) - Frequency of use or execution - Degree of control (involvement of contractors) - Strength/completeness of administrative controls - Sufficiency/formality of training - Other....